## How to review
Instrument capabilities: review copy
Generated from data/annotations.json and data/techniques.json by scripts/draft_annotations.py export.
Grey columns are Stefan's original dataset. The other columns are draft notes about what this TYPE of instrument typically does; nothing here says an accessory is installed unless 'Installed per the dataset' says so.
Please write corrections in the 'Reviewer comments' column (Instruments) or the 'Comments' column (Techniques). The team updates the JSON files from them; edits elsewhere in this file are overwritten on the next export.
Source: dataset = stated in the dataset; llm_draft = drafted with an LLM, unreviewed.
Confidence: high = standard for this instrument type; medium = depends on configuration; low = only with optional hardware the dataset does not list.
Techniques sheet: techniques marked NOT AVAILABLE are ones no listed instrument performs (e.g. micro-CT); the app uses them to say what is missing.

## Instruments
ID | Instrument | Manufacturer | Model | Notes | Location | Contact | Techniques | What it measures | Destructive? | Must come before | Leads to | Typical samples | Limitations | Accessories to confirm | Installed per the dataset | Source and confidence | Review status | Reviewer comments
EQ01 | Extrusion-based 3D bioprinter | EnvisionTEC | 3D-Bioplotter, 4th generation (Manufacturer series) | Two low-temperature heads; one high-temperature head; one UV-light head (multiple wavelengths) | School of Dentistry – user research lab | Avery Bennett (SOD) | Extrusion-based (bio)printing | nothing (makes or prepares samples) | Extrusion-based (bio)printing: not applicable | Bioink / printing material preparation (always) | Confocal laser scanning microscopy; Uniaxial compression testing; Scanning electron microscopy (environmental / ESEM mode) | hydrogels; bioinks; cell-laden materials; thermoplastic scaffold materials | makes constructs; it does not measure any property; lower resolution than light-based printing; printable materials depend on the print head and temperature | suitable needles / nozzles and cartridges for the material; sterile enclosure for printing with cells | two low-temperature heads, one high-temperature head, one UV-light head | Extrusion-based (bio)printing: dataset, high | draft
EQ02 | DLP 3D printer | EnvisionTEC | Envision One |  | UMKC Robert W. Plaster Free Enterprise and Research Center | Avery Bennett (SOD) | Vat photopolymerization 3D printing (DLP / MSLA) | nothing (makes or prepares samples) | Vat photopolymerization 3D printing (DLP / MSLA): not applicable | none | Uniaxial compression testing; Flexural (three- or four-point bend) testing; Uniaxial tensile testing; Scanning electron microscopy (environmental / ESEM mode) | photopolymer resins; dental and biomedical resins | makes parts; it does not measure any property; limited to photocurable resins that suit this printer | resin compatible with the printer; wash and UV post-curing unit | none | Vat photopolymerization 3D printing (DLP / MSLA): dataset, high | draft
EQ03 | Resin 3D printer (MSLA) | ELEGOO | Saturn 4 Ultra 16K |  | School of Dentistry – user research lab | Morgan Ellis (SSE) | Vat photopolymerization 3D printing (DLP / MSLA) | nothing (makes or prepares samples) | Vat photopolymerization 3D printing (DLP / MSLA): not applicable | none | Uniaxial compression testing; Flexural (three- or four-point bend) testing; Uniaxial tensile testing; Scanning electron microscopy (environmental / ESEM mode) | photopolymer resins | makes parts; it does not measure any property; limited to photocurable resins; resin handling required | resin compatible with the printer; wash and UV post-curing station | none | Vat photopolymerization 3D printing (DLP / MSLA): dataset, high | draft
EQ04 | Encapsulator | Büchi | B-390 |  | School of Dentistry – user research lab | Morgan Ellis (SSE) | Vibrating-nozzle encapsulation (microbeads / microcapsules) | nothing (makes or prepares samples) | Vibrating-nozzle encapsulation (microbeads / microcapsules): not applicable | none | Brightfield light microscopy; Stereomicroscopy; Laser diffraction particle sizing | alginate and other hydrogel solutions; polymer solutions; cell suspensions | makes beads and capsules; it does not measure their size or properties; bead size range depends on the nozzle and the formulation | nozzle of the required diameter; concentric nozzle for core-shell capsules; sterile setup for cell encapsulation | none | Vibrating-nozzle encapsulation (microbeads / microcapsules): dataset, high | draft
EQ05 | Analytical mill | IKA | A11 | Exact accessory configuration not specified. | School of Dentistry – user research lab | Taylor Brooks (SOM) | Milling / grinding to powder | nothing (makes or prepares samples) | Milling / grinding to powder: yes | none | Thermogravimetric analysis (TGA); Differential scanning calorimetry (DSC); Fourier-transform infrared spectroscopy (FTIR); X-ray diffraction (XRD) | brittle polymers; dried biological tissue; plant material; small batches of solids | prepares powders; it does not measure particle size; small batches only; heat from grinding can change sensitive samples unless they are cooled | grinding chamber and blade suited to the material; cooling with liquid nitrogen or dry ice for soft or heat-sensitive samples | none | Milling / grinding to powder: dataset, high | draft
EQ06 | Fluorometer | Invitrogen / Thermo Fisher Scientific | Qubit Flex |  | School of Dentistry – user research lab | Taylor Brooks (SOM) | Fluorometric quantification (dye-based assay) | DNA / RNA concentration; protein concentration | Fluorometric quantification (dye-based assay): yes | Nucleic-acid or protein extraction (always) | none | DNA; RNA; protein extracts | only quantifies what a matching assay kit detects; it does not record full fluorescence spectra; does not check purity or integrity of nucleic acids | assay kit for the analyte (dsDNA, RNA or protein); compatible tubes or strips | none | Fluorometric quantification (dye-based assay): llm_draft, high | draft
EQ07 | Inverted confocal microscope | Leica | SP8 STELLARIS | Inverted configuration per source. | School of Dentistry – Confocal & Multiphoton Core | Casey Morgan (SOP) | Confocal laser scanning microscopy; Live-cell time-lapse imaging | location of fluorescent labels; fluorescence intensity (image-based); cell morphology; tissue morphology (histology); live-cell behaviour over time; internal 3D structure (partly: only for transparent or fluorescent samples, to a depth of roughly 100 µm); porosity and pore size (partly: only near the surface of transparent or fluorescently stained samples) | Confocal laser scanning microscopy: depends; Live-cell time-lapse imaging: no | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | fixed or live cells in dishes; thin tissue sections; fluorescently labeled hydrogels and scaffolds | needs fluorescent contrast; imaging depth limited to roughly 100 µm; cannot see inside opaque samples; inverted stand suits samples in dishes or on coverslips | laser lines and detectors matching the fluorophores; objectives of suitable magnification and working distance; stage-top incubator (temperature, CO2) for live cells | none | Confocal laser scanning microscopy: dataset, high; Live-cell time-lapse imaging: llm_draft, low | draft
EQ08 | Upright confocal / multiphoton microscope | Leica | SP8 | Upright multiphoton configuration per source. | School of Dentistry – Confocal & Multiphoton Core | Casey Morgan (SOP) | Confocal laser scanning microscopy; Multiphoton microscopy | location of fluorescent labels; fluorescence intensity (image-based); cell morphology; tissue morphology (histology); internal 3D structure (partly: only for transparent or fluorescent samples, to a depth of roughly 100 µm); porosity and pore size (partly: only near the surface of transparent or fluorescently stained samples); internal 3D structure (partly: only for fluorescent or SHG-active samples (e.g. collagen); deeper than confocal, up to several hundred µm in tissue, but not through opaque materials); porosity and pore size (partly: only for fluorescent or SHG-active samples, to a limited depth) | Confocal laser scanning microscopy: depends; Multiphoton microscopy: depends | Fluorescent labeling / staining (if the sample is not fluorescent itself); Fluorescent labeling / staining (if the sample is neither fluorescent nor SHG-active) | none | thick tissue; explants; fluorescent or collagen-rich scaffolds; intravital samples | needs fluorescent or SHG contrast; deeper than confocal but still limited to several hundred µm; cannot see through opaque materials; upright stand suits thick samples and dipping objectives | tunable infrared multiphoton laser (dataset lists a multiphoton configuration; confirm wavelengths); non-descanned detectors for deep imaging; water-dipping objective with long working distance | none | Confocal laser scanning microscopy: dataset, high; Multiphoton microscopy: dataset, high | draft
EQ09 | High-content epifluorescence imaging system | Keyence | BZ-X810 |  | School of Dentistry – Confocal & Multiphoton Core | Jordan Mercer (SOD) | Widefield (epi)fluorescence microscopy; Brightfield light microscopy; Phase-contrast microscopy; Live-cell time-lapse imaging | location of fluorescent labels; fluorescence intensity (image-based); cell morphology; cell counts / confluence; tissue morphology (histology); live-cell behaviour over time | Widefield (epi)fluorescence microscopy: depends; Brightfield light microscopy: no; Phase-contrast microscopy: no; Live-cell time-lapse imaging: no | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | cell cultures in multiwell plates; tissue sections on slides; whole-slide scans | widefield: little optical sectioning compared with confocal; not suitable for thick or opaque samples | filter cubes matching the fluorophores; phase-contrast objectives; stage-top incubator for live cells; image analysis / cell counting software module | none | Widefield (epi)fluorescence microscopy: dataset, high; Brightfield light microscopy: llm_draft, high; Phase-contrast microscopy: llm_draft, medium; Live-cell time-lapse imaging: llm_draft, low | draft
EQ10 | Environmental scanning electron microscope (ESEM) | Philips | XL30 | Field-emission environmental SEM per source. | School of Dentistry – Electron Microscopy Core | Jordan Mercer (SOD) | Scanning electron microscopy (environmental / ESEM mode); Energy-dispersive X-ray spectroscopy (EDS / EDX) | surface morphology; nanoparticle size and shape (by imaging); particle size distribution (micrometre to millimetre); elemental composition; internal 3D structure (partly: only on cut or fractured cross-sections, which destroys the sample); porosity and pore size (partly: 2D, from surface or cross-section images) | Scanning electron microscopy (environmental / ESEM mode): depends; Energy-dispersive X-ray spectroscopy (EDS / EDX): no | Sputter coating (conductive coating for SEM) (only for high-vacuum imaging of non-conductive samples; ESEM mode can avoid it) | none | polymers and scaffolds; dental materials; hydrated or uncoated biological specimens; powders and particles; fracture surfaces | images surfaces; internal structure only after cutting or fracturing the sample; sample size limited by the chamber and stage; porosity from images is 2D | EDS detector (not listed in the dataset); gaseous / environmental secondary-electron detector for uncoated samples; Peltier cooling stage for hydrated samples; sample stubs and conductive tape | none | Scanning electron microscopy (environmental / ESEM mode): dataset, high; Energy-dispersive X-ray spectroscopy (EDS / EDX): llm_draft, low | draft
EQ11 | Transmission / scanning transmission electron microscope | Philips | CM12 | Manufacturer Philips verified from independent Philips technical documentation. | School of Dentistry – Electron Microscopy Core | Riley Parker (SSE) | Transmission electron microscopy (TEM); Scanning transmission electron microscopy (STEM) | ultrastructure (nanometre-scale structure in thin sections); nanoparticle size and shape (by imaging); internal 3D structure (partly: 2D views of ultrathin sections only; needs the sample to be sectioned) | Transmission electron microscopy (TEM): yes; Scanning transmission electron microscopy (STEM): yes | Ultramicrotomy (ultrathin and semithin sectioning) (for tissue, cells or bulk polymer samples); Negative staining on a TEM grid (for nanoparticles or macromolecules dispersed on a grid) | none | ultrathin sections of fixed tissue, cells or polymers; nanoparticles on grids; negatively stained macromolecules | sample must be thinner than about 100 nm, so bulk samples must be sectioned; very small field of view; beam can damage sensitive samples | STEM detector (instrument name lists STEM; confirm the detector is working); digital camera; TEM grids and suitable sample holder | none | Transmission electron microscopy (TEM): dataset, high; Scanning transmission electron microscopy (STEM): dataset, medium | draft
EQ12 | Upright optical microscope | Nikon | Eclipse E800 | System 1: brightfield, phase contrast, fluorescence; distinguish from system 2. | School of Dentistry – Light Microscopy Core | Riley Parker (SSE) | Brightfield light microscopy; Phase-contrast microscopy; Widefield (epi)fluorescence microscopy | cell morphology; tissue morphology (histology); cell counts / confluence; location of fluorescent labels; fluorescence intensity (image-based) | Brightfield light microscopy: no; Phase-contrast microscopy: no; Widefield (epi)fluorescence microscopy: depends | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | stained tissue sections; semithin sections; cells on slides or coverslips | resolution limited by visible light (about 0.2 µm); needs thin samples on slides | fluorescence filter cubes matching the fluorophores; digital camera | none | Brightfield light microscopy: dataset, high; Phase-contrast microscopy: dataset, high; Widefield (epi)fluorescence microscopy: dataset, high | draft
EQ13 | Inverted optical microscope | Nikon | Eclipse TE300 |  | School of Dentistry – Light Microscopy Core | Alex Rowan (SOM) | Brightfield light microscopy; Phase-contrast microscopy; Widefield (epi)fluorescence microscopy; Live-cell time-lapse imaging | cell morphology; tissue morphology (histology); cell counts / confluence; location of fluorescent labels; fluorescence intensity (image-based); live-cell behaviour over time | Brightfield light microscopy: no; Phase-contrast microscopy: no; Widefield (epi)fluorescence microscopy: depends; Live-cell time-lapse imaging: no | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | live cell cultures in flasks, dishes and plates | resolution limited by visible light; configuration (phase contrast, fluorescence, camera) is not listed in the dataset | phase-contrast objectives and condenser; epifluorescence illuminator and filter cubes; stage-top incubator for live-cell time-lapse; digital camera | none | Brightfield light microscopy: llm_draft, high; Phase-contrast microscopy: llm_draft, medium; Widefield (epi)fluorescence microscopy: llm_draft, low; Live-cell time-lapse imaging: llm_draft, low | draft
EQ14 | Stereomicroscope | Nikon | SMZ 445 |  | School of Dentistry – Light Microscopy Core | Alex Rowan (SOM) | Stereomicroscopy | macroscopic morphology; surface morphology; particle size distribution (micrometre to millimetre) (partly: rough size of beads or particles above ~50 µm with a calibrated camera) | Stereomicroscopy: no | none | none | whole specimens; teeth and dental restorations; printed parts and scaffolds; fracture surfaces; dissections | low magnification (roughly up to 35x for this class); surface only; measurements need a calibrated camera | ring light or other illumination; camera with calibrated measuring software | none | Stereomicroscopy: dataset, high | draft
EQ15 | Universal mechanical testing system | Instron | 5967 | Specific 5967 model confirmed in conversation; general core page does not identify model. | School of Dentistry – departmental/shared equipment | Jamie Collins (SOP) | Uniaxial tensile testing; Uniaxial compression testing; Flexural (three- or four-point bend) testing; Bond strength testing (shear or micro-tensile) | tensile elastic modulus; tensile strength; elongation at break; compressive modulus (stiffness); compressive strength / yield; flexural modulus; flexural strength; bond / adhesion strength (shear or tensile) | Uniaxial tensile testing: yes; Uniaxial compression testing: depends; Flexural (three- or four-point bend) testing: yes; Bond strength testing (shear or micro-tensile): yes | none | Scanning electron microscopy (environmental / ESEM mode); Stereomicroscopy | polymers; porous scaffolds; dental materials and composites; biomaterials; small metal specimens | low-rate electromechanical frame; not for high-cycle fatigue or impact testing; porous scaffolds give an apparent (structural) modulus; porosity itself is not measured; results depend on specimen geometry and fixtures | load cell with a capacity suited to the expected forces (low capacity for soft or porous samples); compression platens; tensile grips suited to the specimen; three- or four-point bend fixture; shear-bond or micro-tensile jig; extensometer for accurate strain; fluid bath or environmental chamber for testing at 37 °C or hydrated | none | Uniaxial tensile testing: llm_draft, high; Uniaxial compression testing: llm_draft, high; Flexural (three- or four-point bend) testing: llm_draft, high; Bond strength testing (shear or micro-tensile): llm_draft, medium | draft
EQ16 | Ultramicrotome | Leica | EM UC7 |  | School of Dentistry – Electron Microscopy Core | Jamie Collins (SOP) | Ultramicrotomy (ultrathin and semithin sectioning); Cryo-ultramicrotomy (sectioning frozen samples) | nothing (makes or prepares samples) | Ultramicrotomy (ultrathin and semithin sectioning): yes; Cryo-ultramicrotomy (sectioning frozen samples): yes | Fixation and resin embedding (always (unless cryo-sectioning)) | Transmission electron microscopy (TEM); Brightfield light microscopy | resin-embedded biological tissue; resin-embedded cells; polymers | measures nothing itself; prepares sections for TEM or light microscopy; destroys the original sample; needs an embedded block | diamond or glass knife; cryo chamber for frozen sections | none | Ultramicrotomy (ultrathin and semithin sectioning): llm_draft, high; Cryo-ultramicrotomy (sectioning frozen samples): llm_draft, low | draft
EQ17 | Thermogravimetric analyzer (TGA) | TA Instruments | TGA550 |  | School of Pharmacy – Drug Discovery and Metabolomics Core | Samira Hayes (SOD) | Thermogravimetric analysis (TGA) | mass loss vs. temperature; thermal stability / decomposition temperature; moisture and volatile content; filler / inorganic residue content | Thermogravimetric analysis (TGA): yes | none | Differential scanning calorimetry (DSC) | polymers; composites; pharmaceuticals; biomaterials; powders | milligram samples; may not represent heterogeneous materials; does not identify the gases released | sample pans (platinum, ceramic or aluminium); purge gas supply (nitrogen; air or oxygen for burn-off of fillers); evolved-gas analysis coupling (MS or FTIR) to identify released gases | none | Thermogravimetric analysis (TGA): dataset, high | draft
EQ18 | Differential scanning calorimeter (DSC) | TA Instruments | DSC250 |  | School of Pharmacy – Drug Discovery and Metabolomics Core | Samira Hayes (SOD) | Differential scanning calorimetry (DSC) | glass-transition temperature (Tg); melting temperature (Tm); crystallization temperature (Tc); degree of crystallinity; heat of reaction / curing enthalpy | Differential scanning calorimetry (DSC): yes | Thermogravimetric analysis (TGA) (recommended first, to keep the DSC below the decomposition temperature) | none | polymers; pharmaceuticals; resins (cure); biomaterials; proteins (with suitable pans) | milligram samples; no chemical identification; run below the decomposition temperature (check with TGA first) | cooling accessory for sub-ambient temperatures and controlled cooling; suitable pans and lids (e.g. hermetic for volatile or wet samples); pan press | none | Differential scanning calorimetry (DSC): dataset, high | draft
EQ19 | Hybrid rheometer | TA Instruments | HR20 Discovery | Installed geometry/accessories not individually confirmed. | School of Pharmacy – Drug Discovery and Metabolomics Core | Drew Navarro (SOP) | Rotational rheometry; Oscillatory rheometry; Dynamic mechanical analysis (DMA) | viscosity vs. shear rate; viscosity vs. temperature; yield stress; viscoelastic moduli (G', G''); gelation and curing kinetics (gel point, cure time); solid-state viscoelastic properties (E', E'', tan delta); glass-transition temperature (Tg); viscosity at processing (extrusion) shear rates (partly: only up to moderate shear rates; extrusion shear rates usually need a capillary rheometer) | Rotational rheometry: depends; Oscillatory rheometry: depends; Dynamic mechanical analysis (DMA): no | none | none | polymer melts; gels and hydrogels; bioinks; pastes; suspensions | temperature range depends on which temperature accessory is installed; maximum shear rate is below typical extrusion shear rates (a capillary rheometer covers those) | temperature-control accessory (Peltier plate, or environmental test chamber for melts); parallel-plate or cone-and-plate geometry; UV-curing accessory for photocuring kinetics; DMA mode / solid clamps (optional on hybrid rheometers) | none | Rotational rheometry: dataset, high; Oscillatory rheometry: dataset, high; Dynamic mechanical analysis (DMA): llm_draft, low | draft
EQ20 | Particle size / zeta-potential analyzer | Malvern Panalytical | Zetasizer Pro (Red Label) |  | School of Pharmacy – Drug Discovery and Metabolomics Core | Drew Navarro (SOP) | Dynamic light scattering (DLS); Electrophoretic light scattering (zeta potential) | particle size distribution (nanometre to low micrometre, in suspension); zeta potential | Dynamic light scattering (DLS): no; Electrophoretic light scattering (zeta potential): depends | Dispersing the sample in a suitable liquid (always) | none | nanoparticles; liposomes; emulsions; protein solutions; colloids | particles from about 1 nm to a few µm only; larger beads or powders need laser diffraction or imaging; requires dilute, well-dispersed samples; a few large particles can bias results | cuvettes suited to the solvent; folded capillary or dip cell for zeta potential; autotitrator for zeta potential vs. pH | none | Dynamic light scattering (DLS): dataset, high; Electrophoretic light scattering (zeta potential): dataset, high | draft

## Techniques
ID | Technique | Kind | Measures | Partly | Destructive? | Must come before | Leads to | Instruments | Comments
rotational_rheometry | Rotational rheometry | measurement | viscosity vs. shear rate; viscosity vs. temperature; yield stress | viscosity at processing (extrusion) shear rates: only up to moderate shear rates; extrusion shear rates usually need a capillary rheometer | depends (the sample is not damaged chemically, but melts and gels are usually not recovered) | none | none | EQ19
oscillatory_rheometry | Oscillatory rheometry | measurement | viscoelastic moduli (G', G''); gelation and curing kinetics (gel point, cure time); yield stress; viscosity vs. temperature | none | depends (small-amplitude tests are non-destructive, but the loaded sample is usually not recovered) | none | none | EQ19
dynamic_mechanical_analysis | Dynamic mechanical analysis (DMA) | measurement | solid-state viscoelastic properties (E', E'', tan delta); glass-transition temperature (Tg) | none | no | none | none | EQ19
capillary_rheometry | Capillary rheometry | measurement | viscosity at processing (extrusion) shear rates; viscosity vs. shear rate; viscosity vs. temperature | none | yes (the sample is extruded) | none | none | NOT AVAILABLE in the dataset
melt_flow_indexing | Melt flow indexing (MFI / MFR) | measurement | melt flow rate (MFR / MFI) | none | yes (the sample is extruded) | none | none | NOT AVAILABLE in the dataset
thermogravimetric_analysis | Thermogravimetric analysis (TGA) | measurement | mass loss vs. temperature; thermal stability / decomposition temperature; moisture and volatile content; filler / inorganic residue content | none | yes (the tested milligram aliquot is heated and usually decomposed) | none | Differential scanning calorimetry (DSC) | EQ17
differential_scanning_calorimetry | Differential scanning calorimetry (DSC) | measurement | glass-transition temperature (Tg); melting temperature (Tm); crystallization temperature (Tc); degree of crystallinity; heat of reaction / curing enthalpy | none | yes (the tested milligram aliquot is sealed in a pan and heated) | Thermogravimetric analysis (TGA) (recommended first, to keep the DSC below the decomposition temperature) | none | EQ18
gel_permeation_chromatography | Gel permeation / size-exclusion chromatography (GPC/SEC) | measurement | molecular weight distribution | none | yes (the sample is dissolved) | none | none | NOT AVAILABLE in the dataset
ftir_spectroscopy | Fourier-transform infrared spectroscopy (FTIR) | measurement | chemical bonds / functional groups; heat of reaction / curing enthalpy | heat of reaction / curing enthalpy: degree of cure from band intensities, not a heat measurement | no | none | none | NOT AVAILABLE in the dataset
x_ray_diffraction | X-ray diffraction (XRD) | measurement | crystalline phase identification; degree of crystallinity | none | no | none | none | NOT AVAILABLE in the dataset
tensile_testing | Uniaxial tensile testing | measurement | tensile elastic modulus; tensile strength; elongation at break | none | yes (tests to failure destroy the specimen; tests limited to the elastic range may not) | none | Scanning electron microscopy (environmental / ESEM mode) | EQ15
compression_testing | Uniaxial compression testing | measurement | compressive modulus (stiffness); compressive strength / yield | none | depends (testing to failure crushes the specimen; testing only in the linear range for stiffness may leave it usable) | none | none | EQ15
flexural_testing | Flexural (three- or four-point bend) testing | measurement | flexural modulus; flexural strength | none | yes (usually tested to failure) | none | none | EQ15
bond_strength_testing | Bond strength testing (shear or micro-tensile) | measurement | bond / adhesion strength (shear or tensile) | none | yes (the bond is broken) | none | Stereomicroscopy | EQ15
indentation_hardness_testing | Indentation hardness testing (micro- or nanoindentation) | measurement | hardness | none | depends (leaves a small indent) | none | none | NOT AVAILABLE in the dataset
micro_computed_tomography | X-ray micro-computed tomography (micro-CT) | imaging | internal 3D structure; porosity and pore size | none | no | none | none | NOT AVAILABLE in the dataset
mercury_intrusion_porosimetry | Mercury intrusion porosimetry | measurement | porosity and pore size | none | yes (the sample is filled with mercury) | none | none | NOT AVAILABLE in the dataset
gas_adsorption | Gas adsorption (BET surface area) | measurement | specific surface area | porosity and pore size: micro- and mesopores only, not scaffold-scale macropores | no | none | none | NOT AVAILABLE in the dataset
atomic_force_microscopy | Atomic force microscopy (AFM) | imaging | surface roughness / topography (quantitative) | hardness: local stiffness by nanoindentation mode | no | none | none | NOT AVAILABLE in the dataset
contact_angle_goniometry | Contact angle goniometry | measurement | wettability (contact angle) | none | no | none | none | NOT AVAILABLE in the dataset
laser_diffraction | Laser diffraction particle sizing | measurement | particle size distribution (micrometre to millimetre) | none | no | none | none | NOT AVAILABLE in the dataset
confocal_microscopy | Confocal laser scanning microscopy | imaging | location of fluorescent labels; fluorescence intensity (image-based); cell morphology; tissue morphology (histology) | internal 3D structure: only for transparent or fluorescent samples, to a depth of roughly 100 µm; porosity and pore size: only near the surface of transparent or fluorescently stained samples | depends (imaging itself does not damage the sample, but labeling, fixation or clearing may) | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | EQ07; EQ08
multiphoton_microscopy | Multiphoton microscopy | imaging | location of fluorescent labels; cell morphology; tissue morphology (histology) | internal 3D structure: only for fluorescent or SHG-active samples (e.g. collagen); deeper than confocal, up to several hundred µm in tissue, but not through opaque materials; porosity and pore size: only for fluorescent or SHG-active samples, to a limited depth | depends (imaging itself does not damage the sample, but labeling may) | Fluorescent labeling / staining (if the sample is neither fluorescent nor SHG-active) | none | EQ08
widefield_fluorescence_microscopy | Widefield (epi)fluorescence microscopy | imaging | location of fluorescent labels; fluorescence intensity (image-based); cell morphology; cell counts / confluence | none | depends (imaging itself does not damage the sample, but labeling or fixation may) | Fluorescent labeling / staining (if the sample is not fluorescent itself) | none | EQ09; EQ12; EQ13
brightfield_microscopy | Brightfield light microscopy | imaging | cell morphology; tissue morphology (histology); cell counts / confluence | none | no (tissue samples are usually sectioned and stained first) | none | none | EQ09; EQ12; EQ13
phase_contrast_microscopy | Phase-contrast microscopy | imaging | cell morphology; cell counts / confluence | none | no | none | none | EQ09; EQ12; EQ13
live_cell_time_lapse_imaging | Live-cell time-lapse imaging | imaging | live-cell behaviour over time; cell morphology | none | no | none | none | EQ07; EQ09; EQ13
stereomicroscopy | Stereomicroscopy | imaging | macroscopic morphology; surface morphology | particle size distribution (micrometre to millimetre): rough size of beads or particles above ~50 µm with a calibrated camera | no | none | none | EQ14
environmental_sem | Scanning electron microscopy (environmental / ESEM mode) | imaging | surface morphology; nanoparticle size and shape (by imaging); particle size distribution (micrometre to millimetre) | internal 3D structure: only on cut or fractured cross-sections, which destroys the sample; porosity and pore size: 2D, from surface or cross-section images | depends (ESEM mode can image uncoated, hydrated samples; the sample must fit the chamber, and internal views need cutting) | Sputter coating (conductive coating for SEM) (only for high-vacuum imaging of non-conductive samples; ESEM mode can avoid it) | none | EQ10
energy_dispersive_xray_spectroscopy | Energy-dispersive X-ray spectroscopy (EDS / EDX) | measurement | elemental composition | none | no | none | none | EQ10
transmission_electron_microscopy | Transmission electron microscopy (TEM) | imaging | ultrastructure (nanometre-scale structure in thin sections); nanoparticle size and shape (by imaging) | internal 3D structure: 2D views of ultrathin sections only; needs the sample to be sectioned | yes (bulk samples must be fixed, embedded and cut into ultrathin sections; the beam can damage sensitive samples) | Ultramicrotomy (ultrathin and semithin sectioning) (for tissue, cells or bulk polymer samples); Negative staining on a TEM grid (for nanoparticles or macromolecules dispersed on a grid) | none | EQ11
scanning_transmission_electron_microscopy | Scanning transmission electron microscopy (STEM) | imaging | ultrastructure (nanometre-scale structure in thin sections); nanoparticle size and shape (by imaging) | none | yes (same sample preparation as TEM) | Ultramicrotomy (ultrathin and semithin sectioning) (for tissue, cells or bulk polymer samples) | none | EQ11
dynamic_light_scattering | Dynamic light scattering (DLS) | measurement | particle size distribution (nanometre to low micrometre, in suspension) | none | no (the dispersed sample can usually be recovered) | Dispersing the sample in a suitable liquid (always) | none | EQ20
electrophoretic_light_scattering | Electrophoretic light scattering (zeta potential) | measurement | zeta potential | none | depends (the applied field can alter sensitive samples such as proteins) | Dispersing the sample in a suitable liquid (always) | none | EQ20
fluorometric_quantification | Fluorometric quantification (dye-based assay) | measurement | DNA / RNA concentration; protein concentration | none | yes (a small aliquot is mixed with dye and not recovered; the rest of the sample is unaffected) | Nucleic-acid or protein extraction (always) | none | EQ06
ultramicrotomy | Ultramicrotomy (ultrathin and semithin sectioning) | sample_preparation | none | none | yes (the block is cut into sections) | Fixation and resin embedding (always (unless cryo-sectioning)) | Transmission electron microscopy (TEM); Brightfield light microscopy | EQ16
cryo_ultramicrotomy | Cryo-ultramicrotomy (sectioning frozen samples) | sample_preparation | none | none | yes (the sample is cut into sections) | none | Transmission electron microscopy (TEM) | EQ16
milling | Milling / grinding to powder | sample_preparation | none | none | yes (the sample is ground to powder) | none | Thermogravimetric analysis (TGA); Differential scanning calorimetry (DSC); Fourier-transform infrared spectroscopy (FTIR); X-ray diffraction (XRD) | EQ05
sputter_coating | Sputter coating (conductive coating for SEM) | sample_preparation | none | none | yes (a thin metal or carbon layer is applied) | none | Scanning electron microscopy (environmental / ESEM mode) | NOT AVAILABLE in the dataset
extrusion_bioprinting | Extrusion-based (bio)printing | fabrication | none | none | not applicable | Bioink / printing material preparation (always) | Confocal laser scanning microscopy; Uniaxial compression testing; Scanning electron microscopy (environmental / ESEM mode) | EQ01
vat_photopolymerization_printing | Vat photopolymerization 3D printing (DLP / MSLA) | fabrication | none | none | not applicable | none | Uniaxial compression testing; Flexural (three- or four-point bend) testing; Uniaxial tensile testing; Scanning electron microscopy (environmental / ESEM mode) | EQ02; EQ03
vibrating_nozzle_encapsulation | Vibrating-nozzle encapsulation (microbeads / microcapsules) | fabrication | none | none | not applicable | none | Brightfield light microscopy; Stereomicroscopy; Laser diffraction particle sizing | EQ04
fixation_and_resin_embedding | Fixation and resin embedding | bench_step | none | none | yes (the sample is chemically fixed and infiltrated with resin) | none | Ultramicrotomy (ultrathin and semithin sectioning) | NOT AVAILABLE in the dataset
fluorescent_labeling | Fluorescent labeling / staining | bench_step | none | none | depends (many protocols fix or permeabilize the sample) | none | none | NOT AVAILABLE in the dataset
negative_staining | Negative staining on a TEM grid | bench_step | none | none | yes | none | Transmission electron microscopy (TEM) | NOT AVAILABLE in the dataset
sample_dispersion | Dispersing the sample in a suitable liquid | bench_step | none | none | depends | none | none | NOT AVAILABLE in the dataset
nucleic_acid_or_protein_extraction | Nucleic-acid or protein extraction | bench_step | none | none | yes | none | Fluorometric quantification (dye-based assay) | NOT AVAILABLE in the dataset
bioink_preparation | Bioink / printing material preparation | bench_step | none | none | not applicable | none | Extrusion-based (bio)printing | NOT AVAILABLE in the dataset